# Introduction

The carbonaceous fraction of the aerosols is by far the least understood and the most difficult to characterize and quantify of the ambient aerosol constituents.

The aerosols total content of carbonaceous material (TC) can be divided into; elemental carbon (EC), inorganic carbon (IC) and organic carbon (OC), of which EC and OC are defined according to the analytical method used for analysis. Quite often the terms EC and black carbon (BC) are mixed. BC is the term used when applying optical methods to quantify the light-absorbing fraction of the aerosols, although this is not specific for carbon. The term EC is based on thermal methods that discriminate between EC and OC.

EC originates from incomplete combustion, and can be both anthropogenic (combustion of fossil fuel) and natural (wild fires). IC can be found in crustal material. OC originates from a wide range of sources, both anthropogenic and natural, and can be both of primary and secondary origin. In the following text, only EC and OC is addressed, while IC is left out, mainly because it typically is present in minor concentrations in ambient aerosols.

While various analytical methods show good agreement with respect to the aerosols content of TC (Schmid et al., 2001), larger uncertainties are associated with the sub fractions EC and OC. This can be attributed to the severe artefacts associated with sampling of OC and the analytical difficulties of separating EC from OC.

Collecting ambient aerosols for the purpose of quantifying OC is challenging because of the need to separate particulate OC from vaporous OC without influencing the gas/particle distribution for each of the compounds constituting the OC fraction. For thermal analysis of OC (and EC), the aerosols need to be collected on Quartz fibre filters to withstand the high temperature (~ 900 °C) attained during analysis. Unfortunately, Quartz fibre filters have a large specific surface area upon which organic vapours adsorb. Further, it is speculated that heating the filters to reduce blank levels might enhance their capacity to adsorb organic vapours.

During sampling, organic vapour phase components may adsorb onto the filter or on the deposited particulate matter impacted on the filter. This is commonly known as the positive artefact, as the amount of carbonaceous material on the filter increases although no particulate matter has been added. Volatile OC that is condensed onto particles trapped by the filter may evaporate during continued sampling, leading to a loss of carbonaceous material that was initially associated with the particles. This loss is denoted as the negative artefact. Particulate matter may also react with vaporous compounds that pass through the filter during sampling, which might change the chemical composition of the deposited material. This reaction may result in both positive and negative artefacts.

Several studies have recognized that the positive artefact is dependent upon the filter face velocity (the volumetric flow rate divided by the exposed filter area; cm s-1) and that the positive artefact decreases as the filter face velocity increases (McDow and Huntzicker, 1990; Turpin et al., 1994). EC, which is not subject to condensation and volatilisation, exhibit no filter face velocity dependence (Turpin et al., 2000).

The Quartz fibre filters content of TC can be analysed using thermal evolution techniques. The TC content is desorbed and converted to either CO2 or CH4, which is quantified by a non-dispersive infrared detector (NDIR) and a flame ionization detector, respectively. Separating TC into EC and OC is more challenging, but can be done by temperature-controlled desorbation/pyrolysis. Various temperature programmes are in use, which can lead to considerable uncertainties with respect to the split between EC and OC.

Currently the most accepted methods for analysing EC and OC are the thermal-optical methods. Such methods account for charring of OC, which can take place when heating of OC is performed in a non-oxidizing environment. If not accounted for, this artificial, pyrolytically generated EC will positively bias the true EC of the sample. This pyrolysis correction is made by continuously monitoring the reflectance and/or the transmittance of the filter, which is largely dominated by EC, and to some extent light absorbing organic carbon. As charring takes place, the reflectance/transmittance drops, whereas it increases when EC/light absorbing organic carbon is oxidized. Hence, the correction is made by measuring the fraction of EC oxidized that is necessary to return the reflectance/transmittance back to the initial value before charring started. This approach assumes that the light extinction per unit mass of pyrolytically generated EC is the same as the light extinction per unit mass of true EC removed until the reflectance/transmittance regains its initial value. Such an assumption is unlikely to be met, thus there is an inherent bias in the split point between EC and OC. The magnitude of this effect cannot be estimated, still it is much less than having no optical correction at all (Gelencsér, 2005).

Currently there is no consensus within the scientific community on how to sample OC in order to account for the positive and the negative artefacts, neither which analytical method should be applied to quantify EC and OC. However, work is going on within the EU-funded project EUSAAR (European Supersites for Atmospheric Aerosol Research), which main objective is to develop, validate, and establish a unified protocol for sampling and analysis of EC and OC in the ambient aerosol. The project aims to get this protocol adapted by both the EMEP task force on measurements and modelling and the WMO Global Atmospheric Watch scientific advisory group.

While awaiting this protocol, EMEP needs to have a set of interim recommendations. 
Such recommendations were put forward during the TFMM meeting in Helsinki (April-2006) and should be regarded as a first step towards the definition of a standardised method for the measurement of carbonaceous aerosol within the EMEP network. These recommendations are based on a literature survey and on the experience we gained during the last years. They aim towards getting a better grip on the quality of the carbonaceous aerosol data gathered within EMEP, and to ensure a certain level of comparability among sites. These recommendations include:

- Sampling time: 24 hrs

The sampling time is set to avoid, as far as possible, considerable changes in weather and air chemistry conditions. A shorter sampling time than 24 h would be preferable in this aspect, but would lead to a too high influence of the field blanks, in particular for those areas with a low PM OC loading. In addition, 24 hours sampling time is consistent with other EMEP filter sampling activity and would more easily allow for comparisons.

- Sampling frequency: Daily

Daily sampling is recommended with respect to improve the time resolution of the data set. However, care should be taken so that not a significant number of samples end up below the detection limit. Hence, the choice of flow rate and filter face velocity should be judged carefully, as well as whether a high or a low PM OC loading can be expected.

# Sampling Equipment: Standard EN 12341

Aerosol samples collected for subsequent analysis of the aerosols content of elemental carbon (EC) and organic carbon (OC) should be collected using samplers operating according to the standard EN 12341 (CEN, 1998), which is the reference method defined in EU Directive 1999/30/EC. Based on experiences from the EMEP EC/OC campaign, the filter face velocity of the sampler should not be lower than 22 cm s-1, in order to minimize influence of fieldblanks; a filter face velocity (volumetric flow rate divided by exposed filter area; cm s-1) in excess of 40 cm s-1 would be preferable, in particular for areas with a low particulate OC loading. In addition, the positive artefact decreases as the filter face velocity increases. On the other hand, negative artefacts will probably increase with increasing filter face velocity due to an increased pressure drop.

- Sampling substrate: Quartz fibre filters

For thermal analysis of EC and OC, the aerosols need to be collected on Quartz fibre filters, as this filter quality withstands the high temperature (~ 900 °C) attained during the analysis. Hence, Teflon-coated Quartz fibre filters and glass fibre filters cannot be used.

- Preheating of Quartz fibre filters:

Commercially available Quartz fibre filters typically contain various amounts of OC, which needs to be removed before sampling takes place. To remove this impurity, filters should be prefired at 850 C for 3 hours.

- Storing of Quartz fibre filters:

After the Quartz fibre filters have been prefired, and prior to being exposed, they should be stored at minimum minus 18 C. Exposed filters should be stored at minimum minus 18C until they are analysed. Non-exposed (prefired) and exposed filters should be stored in petrislides (e.g., Millipore/Whatman) or in prefired (< 500 C for 3 hours) aluminium foil.

- Handling of Quartz fibre filters

Quartz fibre filters are fragile, and fibres and fragments can easily be lost. This problem is enhanced when the filters are prefired, thus the filters should be handled with great care. All handling should be performed using a pair of tweezers; Metal tweezers are preferred. Under no circumstances should you touch the filters with your fingers.

- Determination of PM mass, EC and OC from the same Quartz fibre filter

If the Quartz fibre filters are used both for EC/OC analysis and gravimetric measurements, the filters should be conditioned at 20 °C  1 and 50 % RH  5 for 48 hours according to the procedure for gravimetric determination of PM, as described in section 3.15 of the EMEP manual. The filters should be stored in a freezer at minus 18 °C as soon as possible after the PM mass on the filter has been measured, in order to minimize the possibility of evaporation of vaporous material from the filter.

- Correction for positive artefacts: Quartz fibre filter-behind-Quartz fibre filter (QBQ) (except if denuder is available)

QBQ sampling technique has proved to lead to comparable results as sampling techniques including denuders (Mader et al., 2003; Viana et al., 2006). This is the only simple technique to address sampling artefacts currently available, which only make use of one sampling instrument.

- QBQ – Arrangement of filters

Two Quartz fibre filters are placed directly on top of each other and inserted into the sampler. The top filter is denoted “front filter” whereas the second one simply is called “backup filter”.

- Calculation of the positive sampling artefact of OC using QBQ – Rationale

Based on single filter sampling it is not analytically possible to distinguish between vaporous OC (positive artefact) and particulate OC.

The QBQ approach is a common and simple approach to correct for the positive sampling artifact of OC, using a Quartz fibre filter downstream of a Quartz fibre front filter (McDow and Huntzicker, 1990). According to theory, OC on the Quartz fibre backup filter originates from the vapour phase, as the front filter essentially removes 100 % of the particulate matter. Thus, the backup filter experiences particle free air only.

The QBQ approach is based on an equal adsorption of organic vapours on both the front Quartz fibre filter that collects particles and adsorbs organic vapour and the Quartz fibre backup filter that only adsorbs organic vapours. When subtracting the OC concentration on the Quartz fibre backup filter from the OC concentration obtained from the Quartz fibre front filter the difference will be the particulate OC fraction (See equation 1). To use this type of correction the Quartz fibre front filter and the Quartz fibre backup filter have to be from the same lot to ensure equal adsorptive capacity (Kirchstetter et al., 2001).

OCFront filter - OCBackup filter = Particulate OC		(eq. 1)

When there is not an established equilibrium between the incoming vapour phase condensable organics and the adsorbed phase on the front filter, the front filter will reduce the concentration reaching the backup filter. Hence, it is crucial not to end the sampling too early as this could lead to an underestimation of the positive artefact, since the front filter tend to reach equilibrium first.

It should be emphasized that calculating the particulate organic carbon concentration by the tandem filter approach provides a lower limit estimate of particulate OC, as the negative artefact is not assessed. Neither is the fact that organic compounds may volatilize from the front filter and adsorb onto the quartz backup filter accounted for. However, this is not expected to contribute significantly (McDow and Huntzicker, 1990). Finally, the lack of EC on backup filters rules out the possibility of penetration of particles from the front filter, at least for particles within the same size range as EC.

- Daily calibration of TC with external (traceable) standards.

Checking the instrument calibration with external standards, carbon-containing water-solution, is a common practise, which should be applied when quantifying levels of EC and OC in ambient aerosol samples within EMEP. The concentration of the external standards should reflect the concentration range of the samples content of carbon.

Each filter punch should be spiked with a small volume (10 - 20 µl) of the carbon-containing standard to be sure that the entire water volume is absorbed into the punch and no solution is lost.

The standard could be made from any 100 % water-soluble compound, e.g. sucrose, oxalate. The benefit of using such compounds is that they are naturally present in ambient aerosols.

- EC/OC split quality assurance using standards (pure organics and organic mixtures should be detected as 100 % OC, pure EC should be detected as > 95 % EC).

This is an essential requirement for defining an acceptable method for EC and OC analysis. However, there is no guarantee that methods fulfilling these requirements would lead to comparable results when applied to atmospheric samples.

Recognition of pure organics could be tested from the external standard. However, we are currently not in the position to recommend anything with respect to organic mixtures as there is no scientific basis to base such recommendations on yet.

A standard procedure on how to generate pure EC can be found at http://www.wrapair.org/APACE/OCEC/presentations.htm in the following presentation: “Charring Minimization in Themal Analysis of Aerosol Carbon” & Uncertianties in Optical Charring Correction Schemes” made by Jianzhen Yu, Hong Kong University.

- Every year comparison of the EC/OC split among laboratories reporting EC and OC data to EMEP, based on test filter samples distributed by EMEP-CCC NILU.

Punches of an exposed Quartz fibre filter (8” x 11”) will be distributed among those laboratories that report levels of EC and OC to EMEP. Despite that the “true” level of EC and OC in these samples are not known, it will provide an understanding of how much of the differences observed among the sites that can be attributed to instrumental discrepancies. This would simply be an extension of the standard exchange practise taking place within EMEP, now also to include the analysis of carbonaceous aerosols.

- Field blanks:

One field blank pr week is recommended. The field blanks are to be prefired, handled and transported the same way as the filters to be exposed. At the sampling site, the field blanks should be inserted into the sampler and run for 1 minute, taken out and stored in the transport container in a shelter for the sampling period, and than brought back with the exposed samplers and stored according to the procedure.

When performing sampling according the QBQ - approach, the QBT-approach, or when using a backup filter and a denuder to account for both the positive and the negative artefact, the field blank content should not be subtracted from that of the front filter. The field blank should be subtracted from the front filter only when sampling according to a single filter approach or when using the combination of a denuder and a single filter.

It is recommended to analyse the filters within a month after they have been sampled, in order to minimize artefact due to evaporation or adsorption.